## Slide 1: Erythrocyte sedimentation rate (ESR)

## Slide 2: Introduction and principle
ESR is the rate of RBCs' settle down in millimeters per one hour. Anticoagulated blood is allowed to stand in a special narrow vertical tube (Westergren tube) for one hour. RBCs by time will settle down (sediment) leaving clear plasma above.
Note: normally RBCs sediment by rouleaux formation (RBCs aggregate one on the top of the other forming a stack or column of RBCs). Rouleaux formation happens faster during inflammatory processes due to formation of high proportion of fibrinogen, globulins, and other acute phase reactant proteins which hasten the RBCs sedimentation rate.

## Slide 3
Normal values of ESR:
Adult men 0-17 mm/ hr
Adult women 0- 25 mm/ hr
Extreme elevation if ESR > 100 mm/hr

## Slide 4: Material and instrument:-
1-Westergren pipette (30 cm in length & 2.55 mm in internal width): it is graduated and open at both ends the graduation is from 0 – 200 mm.
2-Westergren pipette rack equipped with leveling screws.
3-3.8% Sodium citrate as an anticoagulant in a ratio of 1:4 with blood.
4-Syringe for the withdrawal of blood from the vein of the patient.

## Slide 5: Procedure:-
1-Withdraw 2 cc of blood from the patients vein using the syringe
2-Put 0.4 cc of sodium citrate in a plain test tube
3-Immediately add 1.6 cc blood from the syringe to the plain tube and shake it for mixing
4-the Westergren pipette to exactly the 0 mark making certain that there are no air bubbles at all in the blood
5-Place the pipette vertically on the rack and leave it for one hour
6-At the end of the 60 minutes read the number of millimeters at graduation scale the top of blood column have reached (i.e.: the height of clear plasma above the upper limit of the column of sedimentary cells) the result is the ESR in mm/ in 1 hour.

## Slide 6: Medical considerations:
There are generally two tests to monitor the inflammatory process during disease state, ESR and C-reactive protein (CRP), so that they are nonspecific markers of inflammation. Generally, ESR does not change as rapidly as does CRP. CRP is not affected by as many other factors as is ESR, making it a better marker of

## Slide 7
inflammation. However, because ESR test is easily performed, ESR is an initial test when one needs to know if there is an inflammation.
Note: ESR increased is typically a result of globulins or fibrinogens (serum protein electrophoresis to determine which of them or both is causing the elevated ESR).

## Slide 8: increased ESR occurs in the following conditions
Physiological: pregnancy, aging,
Pathological
anemia
Temporal arteritis
polymyalgia rheumatica
Leukemia and Lymphoma (immunoglobulins are secreted in high amounts
Children with rheumatoid arthritis or Kawasaki’s Disease.
Drugs: e.g. (methyldopa, oral contraceptives, theophylline, and vitamin A
Infection (Tuberculosis ) may cause extreme elevation.
Renal disease may cause extreme elevation.
Collagen vascular disease (Rheumatoid arthritis) may cause extreme elevation.
Multiple Myeloma or some other type of dysproteinemia may cause extreme elevation.
Metastatic Malignant tumors especially if metastasis occur may cause extreme elevation.
Rouleaux Formation: RBCs have stacked together into long chains
Nonspecific cause (2% of cases of high ESR).

## Slide 9: Decreased ESR occurs in:
Congestive heart failure
Polycythemia
Hereditary spherocytosis
Drugs: albumin, aspirin, and cortisone.
Erroneous values (laboratory errors)
Artifactual (False) high ESR due to:
Tilted ESR tube
High temperature
Sample dilution
Artifactual (False) low ESR due to:
Inadequate anticoagulation with clotting of the blood sample (consequently will consume fibrinogen)
Short ESR tube
Vibration during testing

## Slide 10: Rouleaux Formation: RBCs have stacked together into long chains