Introduction

In modern history, metals have been used as implants since more than 100 years ago when "Lane" first introduced metal plate for bone fracture fixation in 1895. The first metal alloy developed specifically for human use was the “vanadium steel” which was used to manufacture bone fracture plates (Sherman plates) and screws. But it was no longer used in implants because its corrosion resistance is inadequate in vivo environment.

Advantages and disadvantages of metallic Biomaterials:

Advantages

- High strength.

- High hardness.

- Fatigue and impact resistance.

- Wear resistance.

- Easy fabrication.

- Easy to sterilize.

- Shape memory.

- inert

Disadvantages

- High modulus.

- High corrosion.

- Metal ion sensitivity and toxicity.

- High density.

In general metallic biomaterials can be grouped in the following categories:

  - Stainless steels.

  - Co-based alloys.

  - Titanium-based alloys.

- Metals and their alloys are widely used as biomedical materials. On one hand, metallic biomaterials cannot be replaced by ceramics or polymers at present because the  mechanical strength and toughness are the most important safety requirements for a biomaterial under load-bearing conditions. On the other hand, metallic materials sometimes show toxicity and are fractured because of their corrosion and mechanical damages . Therefore, development of new alloys is continuously trialed.

STAINLESS STEELS

is stronger and more resistant to corrosion than the vanadium steel.

TABLE (1): Compositions of 316L Stainless Steel (American Society for

Testing and Materials, F139—86. p. 61. 1992).

| Element | Composition (%) |
| --- | --- |
| Carbon | 0.03 max. |
| Manganese | 2.00 max. |
| Phosphorus | 0.03 max. |
| Sulfur | 0.03 max. |
| Silicon | 0.75 max. |
| Chromium | 17.00 – 20.00 |
| Nickel | 12.00 – 14.00 |
| Molybdenum | 2.00 -4.00 |

The Ni stabilizes the austenitic phase () face centered cubic crystal (FCC) structure at room temperature and enhances corrosion resistance.

The advantages of stainless steels, especially type 316 and 316L over other grades of steel:

- Biocompatible.

- This group of stainless steels is nonmagnetic and possesses better corrosion resistance than any others.

TABLE (2): Mechanical Properties of 316L Stainless Steel for Implants (American Society for Testing and Materials. F139—86, p.61, 1992).

A wide range of properties exists depending on the heat treatment (annealing to obtain softer materials) or cold working (for greater strength and hardness).

(Co Cr) ALLOYS

The castable CoCrMo alloy.

The castable CoCrMo alloy has been used for many decades in dentistry and, relatively recently, in making artificial joints.

TABLE (3): Mechanical Property Requirements of Co-Cr Alloys

The corrosion products of CoCrMo are more toxic than those of stainless steel 316L.

The advantages and disadvantage of CoCrMo:

 wrought or forged forms has the highest strength/wear resistance,

 hardest to fabricate,

 may produce cobalt or chromium ion sensitivity/toxicity.

Ti ALLOYS

Ti-based alloys are finding ever-increasing applications in biomaterials due to their excellent mechanical, physical and biological performance.

- Pure Ti and Ti6Al4V

Commercially pure titanium (Ti CP) and Ti-6Al-4V are the two most common titanium base implant biomaterials. These materials are classified as biologically inert biomaterials.

Comparison betweenTitanium alloy (Ti-6Al-4V) versus Titanium (Ti) metal

- titanium alloy is stronger than titanium metal,

-both have the best corrosion resistance,

-both have excellent bone bonding.

The modulus of elasticity of these materials is about (110 GPa.) This is much lower than stainless steels and Co-base alloys modulus (210 and 240 GPa.) respectively.

2- TiNi Alloys

The titanium—nickel alloys show unusual properties, that is, after it is deformed the material can snap back to its previous shape by following heating of the material. This phenomenon is called shape memory effect (SME).

(Nitinol) exhibits an exceptional SME near room temperature, if it is plastically deform below the transformation temperature, it revert back to its original shape as the temperature is raised.

In order to develop such devices, it is necessary to understand fully the mechanical and thermal behavior associated with the martensite phase transformation. A widely known NiTi alloy is (55-Nitinol) (55 wt% atomic Ni)

Some possible applications of shape memory effects alloys are orthodontic dental archwires, intracranial aneurysm, skull clip, contractible muscles for artificial heart, vascular stent, and catheter wire guide.

In general , The advantage and disadvantage of Ti alloys are :

Advantages

•Easily formed.

•Highly biocompatible.

•Outstanding corrosion resistance.

•Better than stainless steel and cobalt-chromium alloys.

•Forms protective oxide (TiO2) layer.

•Low elastic modulus.

Disadvantages

•Poor wear resistance.

•Should not be used in articulated surfaces such as hip or knee joints unless surface-treated through ion implantation which improves wear resistance.

In General Titanium alloys used in implants present three main problems:

- High cost because the amount of processing energy and melting and casting difficulties.

- Higher elastic modulus compared to bone.

- Although the inert behavior of Ti is a good property, its bone attachment is difficult because it do not react with the human tissues.

The following table represents the comparison of mechanical properties of metallic biomaterials with bone.

Table (4): Comparison of mechanical properties of metallic biomaterial with bone.

| Material | Young’s Modulus, E (GPa) | Yield Strength, σy (MPa) | Tensile Strength, σUTS (MPa) | Fatigue Limit, σend (MPa) |
| --- | --- | --- | --- | --- |
| Stainless steel | 190 | 221–1213 | 586–1351 | 241–820 |
| Co-Cr alloys | 210–253 | 448–1606 | 655–1896 | 207–950 |
| Titanium (Ti) | 110 | 485 | 760 | 300 |
| Ti-6Al-4V | 116 | 896–1034 | 965–1103 | 620 |
| Cortical bone | 15–30 | 30–70 | 70–150 |  |

Bio metals applications

Table (3): Implants division and type of metals used

| Division | Example of implants | Type of metal |
| --- | --- | --- |
| Cardiovascular | Stent Artificial valve | 316L SS; CoCrMo; Ti Ti6Al4V |
| Orthopedics | Bone fixation (plate, screw, pin) Artificial joints | 316L SS; Ti; Ti6Al4V CoCrMo; Ti6Al4V; Ti6Al7Nb |
| Dentistry | Orthodontic wire Filling | 316L SS; CoCrMo; TiNi; TiMo AgSn(Cu) amalgam, Au |
| Craniofacial | Plate and screw | 316L SS; CoCrMo; Ti; Ti6Al4V |
| Otorhinology | Artificial eardrum | 316L SS |

Failure of metals for biomedical devices

- Corrosion

Metal implant is prone to corrosion during its services due to corrosive medium of implantation site. Types of corrosion that frequently found in implant applications are fretting, pitting and fatigue.

- Fatigue and fracture

During its service most of metallic implants are subjected to cyclic loading inside the human body which leads to the possibility for fatigue fracture. Factors determine the fatigue behavior of implant materials include microstructure of the implant materials.

- Wear

Together with corrosion process, wear is among the surface degradation that limits the use of metallic implant such as Ti alloy. Removal of dense oxide film which naturally formed on the surface of this metallic implant in turn caused wear process. In fact, the major factor that causing premature failure of hip prostheses is due to the wear process with multiple variables interact and thus increase the resultant wear rates.