## Slide 1: Nephelometry and Turbidimetry

## Slide 2: Turbidimetry
• Turbidimetry is involved with measuring the amount of transmitted light (and calculating the absorbed light) by particles in suspension to determine the concentration of the substance in question.
Amount of absorbed light, and therefore, concentration is dependent on ; 1) number
of particles, and 2) size of particles.
• Measurements are made using light spectrophotometers

## Slide 3: Clinical Applications
• Determination of the concentration of total protein in biological fluids such as urine and CSF which contain small quantities of protein (mg/L quantities) using trichloroacetic acid
• Determination of amylase activity using starch as substrate. The decrease in turbidity is directly proportional to amylase activity.
• Determination of lipase activity using triglycerides as substrate. The decrease in turbidity is directly proportional to lipase activity.

## Slide 4: Nephelometry. Principle
• Nephelometry is concerned with measurement of scattered light from a cuvette containing suspended particles in a solution.
• The components of a nephelometer are the same as a light spectrophotometer except that the detector is placed at a specific angle from the incident light.
• The detector is a photomultiplier tube placed at a position to detect forward scattered light.
Detectors may be placed at 90o, 70o or 37o depending on the angle at which most scattered light are found.

## Slide 5: principle
Since the amount of scattered light is far greater than the transmitted light in a turbid suspension, nephelometry offers higher sensitivity than turbidimetry.
The amount of scattered light depends on the size and number of particles in suspension.
For most clinical applications, the light source is a tungsten lamp giving light in the visible region
For higher sensitivity and for applications that determine the size and number of particles in suspension, laser light nephelometers is used.

## Slide 6: Clinical applications of nephelometry.
• Widely used to determine concentrations of unknowns where there is antigen-antibody reactions such as Determination of immunoglobulins (total, IgG, IgE, IgM, IgA) in serum and other biological fluids
o Determination of the concentrations of individual serum proteins; hemoglobin, haptoglobin, transferring, c-reactive protein, a,1-antitrypsin, albumin (using antibodies specific for each protein)
o Determination of the size and number of particles (laser-nephelometr}

## Slide 7: Considerations in turbidimetry and nephelometry
The reaction in turbidimetry & nephelometry does not follow Beer's Law
• Therefore, standard curves must be plotted and the concentration of the unknown is determined from the standard curve.
• Because the absorbance is dependent on both number and size of particles, the standard solution which is used for the standard curve must have similar size in suspension as unknown.
• Because some precipitation and settlement of particles may occur with time, in order to obtain good accuracy it is important to ; a) mix the sample well prior to placing the cuvette in the instrument, and, b) keep the same time for measurement of every sample throughout the measurement.

## Slide 8
Kinetic reactions (measurement of the progress of reaction with time) provides higher degree of accuracy, sensitivity, precision and less time than end-point reactions (measuring the reaction at the start and finish of the reaction)
o Additionally in kinetic reactions there is no need for reagent blank since the previous reading is taken as the base-line for the next reading.
o Kinetic reaction may be taken in 60, 90 or 120 seconds (taking readings at 10 seconds intervals), whereas endpoint reactions may take much longer time e.g. 15 -120 minutes.

## Slide 9: Selection of a wavelength
• If both solution and suspended particles are colorless, then use any wave length in the visible range
• If the solution is coloured but the particles are not coloured, then use a wave length that gives minimum absorption for the solution
• If the particles are coloured and the solution is colorless then use a wavelength that gives maximum absorption with the particles
• If both solution and particles are coloured then use two wavelengths; one that gives minimum absorbance for the solution and the other one maximum absorbance for the particles. Subtract the solution absorbance from the particles absorbance.

## Slide 10: Turbidimetry and Nephelometry
When particles are suspended in a solution in a cuvette, they make the solution unclear (turbid).
Incident light entering the cuvette will be subjected to three reactions;
1- some of the light will be absorbed (blocked) by the particles
2- some will be transmitted through the cuvette
3- some will be scattered in various directions.

## Slide 11
Turbid metric + nephlometeric analysis:
When part of incident radiant energy is dissipated by absorption, reflection, and refraction, while the remainder is transmit light as a function of the concentration of the dispersed phase is the basis of turbid metric analysis

## Slide 13
 | Nephlometer | Turbidimeter
Definition | the measurement of the intensity of scattered light at right angles to the direction of the incident light as a function of the concentration of the dispersed phase ,It is most sensitive for very dilute suspensions (100 mg/L). | Light passing through a medium with dispersed particles, so the intensity of light transmitted is measured.
Instrument used | Nephlometery machine | spectrophotometer
Type of light measured | Scattered light | Transmitted light
Arrangement of photometer | measure the light scattered at right angle to the direction of the propagation of light from the source. It could be movable detectors which allow operator to vary the angle of detection | made in the same direction as the propagation of the light from the source.
Clinical uses | Ag-Ab rxn, immunocomplex rxn,ppts, lipoprotein | Ag-Ab rxn, immunocomplex rxn,ppts, liver dis, protein in urine or CSF

## Slide 14
Nephelometry and Turbidimetry

## Slide 15: Instrumentation:
1- light source:
Tungsten its relatively low intensity makes it less useful for samples with low light scattering. Alternatives are:
quartz halogen lamp, xenon lamp and laser which have higher intensities than tungsten lamp.
2-lens assembly:
Light enter the sample holder through lens assembly.

## Slide 16
3-there is provision for the insertion of filter between the sample and source of light(monochromator).
4- detector (photo –cell):
It is shielded to minimize interference from stray light.
5-Read out device:
Light intensity is converted to an electrical signal by the detector .

## Slide 17: Example:
The turbidity of a dilute barium sulphate suspension.
The concentration of the reactants must be controlled by adding pure solid barium chloride of definite grain size.
-NACL and HCL are added before the precipitation in order to inhibit the growth of microcrystal of barium sulphate
-A glycerol ethanol solution helps to stabilise the turbidity.
The reaction vessels is shaken gently in order to obtain a uniform particle size.
Each vessel should be shaken at the same rate and the same number of times.

## Slide 19: MEASURMENTS:
1-plug the instrument into ground outlet.
2- choose desirable scale from 0-10
starting with the highest conc. (for std 1=scale 10)
3- turn power switch on.
4-selecte desirable range by range selector at desirable position .
5- select filter required.

## Slide 20
6- transfer your standards in the cleaned cell and place them in cell holder.
7- remove the standards.
8- fill the second cell with blank to set zero .
9- check the reading of the standards again.
10- measure your unkown.
11- draw calibration curve.

## Slide 21: Precaution:
-Number and size of the particles should remain constant if repeated preparation are made
-clean cell & filter
- avoid air bubbles (high reading).
-dilute sample if there is need.
-prepare the blank ,standards, Sample at the same time.(to avoid ppt)

## Slide 22: Advantages and disadvantages :
Advantages : rapidity of procedure and simplicity of the measurements.
Disadvantages: lack of accuracy of the measurements.